## Integration
Chromatogram and Results
Injection Details
Injection Name: |  | 14.2.3 |  |  | Run Time (min): | 25
Vial Number: |  | 20 |  |  | Injection Volume: | 5000
Injection Type: |  | Unknown |  |  | Channel: | ECD_1
Calibration Level: |  |  |  |  | Wavelength: | n.a.
Instrument Method: |  | Medmerry1 |  |  | Bandwidth: | n.a.
Processing Method: |  | Medmerry 1 proc |  |  | Dilution Factor: | 1
Injection Date/Time: |  | 2016-12-22 19:35 |  |  | Sample Weight: | 1
Chromatogram
Integration Results
No. | Peak Name | Retention Time | Area | Height | Relative Area | Relative Height | Amount
|  | min | µS*min | µS | % | %
n.a. | Fluoride | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
1 | Chloride | 6.5405 | 20.0597 | 72.3082 | 19.1356 | 37.3509 | 86.7373
n.a. | Nitrite | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | Bromide | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | Nitrate | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
2 | Phosphate | 15.9838 | 23.7407 | 33.9595 | 22.647 | 17.5418 | 512.8981
3 | Sulfate | 18.4472 | 61.0291 | 87.3238 | 58.2174 | 45.1073 | 512.1366
Total: |  |  | 104.8296 | 193.5915 | 100 | 100

## Peak Analysis
Peak Analysis
Injection Details
Injection Name: |  | 14.2.3 |  |  | Run Time (min): | 25
Vial Number: |  | 20 |  |  | Injection Volume: | 5000
Injection Type: |  | Unknown |  |  | Channel: | ECD_1
Calibration Level: |  |  |  |  | Wavelength: | n.a.
Instrument Method: |  | Medmerry1 |  |  | Bandwidth: | n.a.
Processing Method: |  | Medmerry 1 proc |  |  | Dilution Factor: | 1
Injection Date/Time: |  | 2016-12-22 19:35 |  |  | Sample Weight: | 1
Chromatogram
Peak Results
No. | Peak Name | Retention Time | Width (50%) | Type | Resolution (EP) | Asymmetry (EP) | Plates (EP)
|  | min | min
n.a. | Fluoride | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
1 | Chloride | 6.5405 | 0.23029 | M | 12.6576 | 1.4544 | 4469
n.a. | Nitrite | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | Bromide | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | Nitrate | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
2 | Phosphate | 15.9838 | 0.65007 | M | 2.2872 | 1.5255 | 3349
3 | Sulfate | 18.4472 | 0.62083 | M | n.a. | 0.88659 | 4891

## SST
Chromatogram and SST Results
Injection Details
Injection Name: |  | 14.2.3 |  | Run Time (min): | 25
Vial Number: |  | 20 |  | Injection Volume: | 5000
Injection Type: |  | Unknown |  | Channel: | ECD_1
Calibration Level: |  |  |  | Wavelength: | n.a.
Instrument Method: |  | Medmerry1 |  | Bandwidth: | n.a.
Processing Method: |  | Medmerry 1 proc |  | Dilution Factor: | 1
Injection Date/Time: |  | 2016-12-22 19:35 |  | Sample Weight: | 1
Chromatogram
SST Results
No. | Name | Inj.Condition | Peak | Test Result | Injection
n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
| Number of executed test cases: | n.a. | Total Result: | Passed

## Chromatogram
Chromatogram