## Integration
Chromatogram and Results
Injection Details
Injection Name: |  | 20ppm 2 |  |  | Run Time (min): | 25
Vial Number: |  | 34 |  |  | Injection Volume: | 5000
Injection Type: |  | Unknown |  |  | Channel: | ECD_1
Calibration Level: |  |  |  |  | Wavelength: | n.a.
Instrument Method: |  | Medmerry1 |  |  | Bandwidth: | n.a.
Processing Method: |  | Medmerry 1 proc |  |  | Dilution Factor: | 1
Injection Date/Time: |  | 2016-12-23 02:05 |  |  | Sample Weight: | 1
Chromatogram
Integration Results
No. | Peak Name | Retention Time | Area | Height | Relative Area | Relative Height | Amount
|  | min | µS*min | µS | % | %
1 | Fluoride | 3.9207 | 1.0963 | 4.4662 | 7.1438 | 11.2908 | 3.6433
2 | Chloride | 6.4573 | 3.8346 | 13.5942 | 24.9887 | 34.3667 | 16.5807
3 | Nitrite | 8.1807 | 2.2845 | 6.3406 | 14.8873 | 16.0292 | 19.1792
4 |  | 10.2207 | 1.5459 | 3.757 | 10.0738 | 9.4979 | n.a.
n.a. | Bromide | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
5 | Nitrate | 11.9073 | 1.9353 | 4.069 | 12.6116 | 10.2865 | 17.7785
6 |  | 15.704 | 2.0592 | 3.3098 | 13.4187 | 8.3673 | n.a.
n.a. | Phosphate | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
7 | Sulfate | 17.494 | 2.5897 | 4.0196 | 16.876 | 10.1616 | 21.7319
Total: |  |  | 15.3454 | 39.5563 | 100 | 100

## Peak Analysis
Peak Analysis
Injection Details
Injection Name: |  | 20ppm 2 |  |  | Run Time (min): | 25
Vial Number: |  | 34 |  |  | Injection Volume: | 5000
Injection Type: |  | Unknown |  |  | Channel: | ECD_1
Calibration Level: |  |  |  |  | Wavelength: | n.a.
Instrument Method: |  | Medmerry1 |  |  | Bandwidth: | n.a.
Processing Method: |  | Medmerry 1 proc |  |  | Dilution Factor: | 1
Injection Date/Time: |  | 2016-12-23 02:05 |  |  | Sample Weight: | 1
Chromatogram
Peak Results
No. | Peak Name | Retention Time | Width (50%) | Type | Resolution (EP) | Asymmetry (EP) | Plates (EP)
|  | min | min
1 | Fluoride | 3.9207 | 0.19831 | M | 6.8651 | 2.2929 | 2165
2 | Chloride | 6.4573 | 0.2377 | M | 3.7531 | 1.7604 | 4088
3 | Nitrite | 8.1807 | 0.30413 | M | 3.6358 | 1.8304 | 4008
4 |  | 10.2207 | 0.35795 | M | 2.5681 | 1.7217 | 4517
n.a. | Bromide | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
5 | Nitrate | 11.9073 | 0.41705 | M | 4.5248 | 1.6997 | 4516
6 |  | 15.704 | 0.57307 | M | 1.8371 | 1.3675 | 4160
n.a. | Phosphate | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
7 | Sulfate | 17.494 | 0.57667 | M | n.a. | 1.4327 | 5098

## SST
Chromatogram and SST Results
Injection Details
Injection Name: |  | 20ppm 2 |  | Run Time (min): | 25
Vial Number: |  | 34 |  | Injection Volume: | 5000
Injection Type: |  | Unknown |  | Channel: | ECD_1
Calibration Level: |  |  |  | Wavelength: | n.a.
Instrument Method: |  | Medmerry1 |  | Bandwidth: | n.a.
Processing Method: |  | Medmerry 1 proc |  | Dilution Factor: | 1
Injection Date/Time: |  | 2016-12-23 02:05 |  | Sample Weight: | 1
Chromatogram
SST Results
No. | Name | Inj.Condition | Peak | Test Result | Injection
n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
| Number of executed test cases: | n.a. | Total Result: | Passed

## Chromatogram
Chromatogram